Details of the developed LabView programs

- R Test

This program is to perform 4-probe resistance measurement using delta mode with Keithley 6221 and 2182A.

It uses positive and negative current to cancel out any bias to the measurement system.

Finite Count represents the total number of voltage measurements. It has to be an even number (this ensures equal number of positive and negative measurements to cancel out any biasing).

High Level and Low Level must have equal and opposite current values.

- RT Measure

This is a controlled RT measurement program based on delta mode measurement with Keithley 6221 and 2182A.

Start Value and Stop values are assigned with a suitable ramp rate to perform the measurement.

Note: - It uses a ramp rate of 1K/min (by default) to reach to the start value. Once it reaches the start value, then the assigned ramp rate is used.

There is a dual cycle button that can be used to do the forward and backward measurement (heating & cooling). If the dual cycle button is enabled, it will do both the measurements; forward and backward.

- IV Test

This is designed to do IV Test to know the correct range of currents to be used for IV Measurements. At low temperature, it can be used to find the critical current. To save the IV Test values, enable the save button and then run the program, it will pop up a window to save the results. Make sure that the file extension is “.csv”.

- IV Measure

This program is developed to perform IV measurements at different temperatures using differential conductance mode. The fields indicated in blue color are the control parameters. One can set the start value, stop value, step size, ramp rate, voltage compliance, holding time and relaxation time.

Holding time is the time duration for which the setpoint holds for stabilization of the temperature before starting the measurement at the specified temperature.

Relaxation time is the time duration for which the sample is set to relax after a measurement is done. This is introduced to handle the sample heating due to current. For larger currents such as 10mA, heating is too much and therefore, an appropriate relaxation time is needed for smooth measurements.

I-V Meas button must be enabled to start the measurement. Once it is enabled the measurement will start when temperature reaches to the start value. If it is not enabled, the program will wait at the start value.

The fields shown in blue are the control parameters and those shown in green are indicator parameters.

Note: The step size in differential parameters must be chosen such that there are not more than 3000 points in total. For example,

Start value (A) = 0

Stop value (A) = 10mA

Step size (A) = 5uA

This will result in 2000 measurement points and this will work. But if we choose a step size of 2uA, it will result in 5000 measurement points and this gives error because of limited buffer memory of Keithley 2182A

- Heater Control

This program is to control the heater outputs. Both loops can be controlled simultaneously by using their on/off button.

- Cooldown RT

This is a program to do RT measurements without controlling the heater. When system is being cooldown, use this program to track the cooling as well as R-T measurement. It saves all the four temperature sensors data to the file.

Three measurements can be done simultaneously, namely sample-1, sample-2 and sample-3. However, these measurements are not parallel measurements, it will be done in sequential manner one by one in rotation.

Important Note: All the programs by default uses 25ohm heater configuration.  However, the heater installed to sample positioners are of 50ohm. The other three heaters are of 25ohm.

All the programs are developed in LabView 2024Q1, and therefore, LabView runtime 2024Q1 is needed to run the applications on a computer where LabView is not installed.